## General
Monitoring data | Data Type | Logging Required | cRIO | Param | Control Data | Data Type | Logging Required
MFM |  | Yes | NI 9205 - AI0 | 0 to 10 V | Store Preflight Log (array of boolean)
File Transfer to FM Laptop
cRIO rtc status | Boolean
Data Logging status | Boolean
ptp_sync | Boolean
cRIO temperature
Free disk space (byte)
Free memory (byte)
CPU Usage (%)
chk interface instance
cRIO console output | string

## CO
ASRL8::INSTR
Monitoring data | Data Type | Logging Required | cRIO | Params |  | Control Data | Data Type | Logging Required | cRIO
Conc | single_float | Yes |  |  |  | CO CAL | Boolean | Yes
Sig | single_float | Yes |  |  |  | CO Vacuum | Boolean | Yes | NI 9472 - DO5
Sens | single_float | Yes |  |  |  | Cal Record Request | Boolean | No
Zero | single_float | Yes
Flow? L | single_float | Yes
Temp? L | single_float | Yes
Flow? M | single_float | Yes
Temp? M | single_float | Yes
Press? M | single_float | Yes
Press? F | single_float | Yes
Press? C | single_float | Yes
Temp? C | single_float | Yes
Calconc? | single_float | Yes
err | text | Yes
CO Cal count | signed_int | Yes
CO Error Counter | signed_int | Yes
ALEG1_press_delivery | single_float | Yes | NI 9205 - AI1 |  | Sensortechnics model CTU9100
ALREG1_press_content | single_float | Yes | NI 9205 - AI2 |  | Wika model A10
ALREG2_press_content | single_float | Yes | NI 9205 - AI3 |  | Wika model A10
ALREG3_press_content | single_float | Yes | NI 9205 - AI4 |  | Wika model A10
CO vacuum status | signed_int | Yes
In CAL? | signed_int | Yes
Cal Record (date, time, sens, zero) |  | No
Log Data Format
Date & Time
CPU Load
MFM
PTP Sync
CO CAL Request
Flight Number
UDP Pack Num
Conc
Sig
Sens
Zero
Flow? L
Temp? L
Flow? M
Temp? M
Press? M
Press? F
Press? C
Temp? C
Calconc?
err
CO Cal count
CO Error Counter
ALEG1_press_delivery
ALREG1_press_content
ALREG2_press_content
ALREG3_press_content
CO vacuum status
In CAL?

## FGGA
ASRL7::INSTR
Monitoring data | Data Type | Logging Required | Control Data | Data Type | Logging Required | cRIO | Control param
Pressure (Torr) |  | Yes | Flush High Tank (V2, V3, V4) & MFC_1 set from 0 to 3* -> for 15 secs -> all off & set MFC_1 to 0 | Boolean | No | NI 9472 DO1, DO2, DO3
Temp (DegC) |  | Yes | Flush Low Tank (V2, V4) & MFC_1 set from 0 to 3* > for 15 secs -> all off & set MFC_1 to 0 | Boolean | No | NI 9472 DO1, DO3
Fit Flag | integer | Yes | Flush Target Tank (V4) & MFC_1 set from 0 to 3* > for 15 secs -> all off & set MFC_1 to 0 | Boolean | No | NI 9472 DO3
RDA Time (uSec) |  | Yes | High Cal (V2,V3,V4) & MFC_1 set to 3*, on for 45 sec -> with (V1) on for 60 sec -> all off & set MFC_1 to 0 | Boolean | Yes | NI 9472 DO1, DO2, DO3 -> DO0
RDB Time (uSec) |  | Yes | Low Cal (V2,V4) & MFC_1 set to 3*, on for 45 sec -> with (V1) on for 60 sec -> all off & set MFC_1 to 0 | Boolean | Yes | NI 9472 DO1, DO3 -> DO0
CH4 Dry (ppb) |  | Yes | Target Cal (V4) & MFC_1 set to 3*, on for 45 sec -> with (V1) on for 60 sec -> all off & set MFC_1 to 0 | Boolean | Yes | NI 9472 DO3 -> DO0
CO2 Dry (ppm) |  | Yes | Synchronise Span Cal (High Low) with CO | Boolean | Yes
H2O (ppm) |  | Yes | FGGA Pump Auto (check WoW) | Boolean | Yes | NI 9263 - AO0 | 0 off ;1.4V ON
MFC_1 absolute pressure (PSIA) |  |  | FGGA pump Manual | Boolean | Yes | NI 9263 - AO0 | 0 off ;1.4V ON
MFC_1 temperature (degC) |  |  | Restart FGGA | Boolean | Yes
MFC_1 volumetic flow (lpm) |  |  | *CAL MFC_1 SLPM Value (*pick up from the manual setting) | double |  |  | RS232 | ASRL3::INSTR
MFC_1 mass flow (slpm)
MFC_1 set point (slpm) |  |  | Shutdown
V1
V2
V3
V4
Restart FGGA
FGGA Pump | double
CAL MFC_1 Set Value | double
Log Data Format
Date & Time
CPU Load
MFM
PTP Sync
Flight Number
Pressure (Torr)
Temp (DegC)
Fit Flag
RDA Time (uSec)
RDB Time (uSec)
CH4 (ppb)
CO2 (ppm)
H2O (ppm)
MFC_1 absolute pressure (PSIA)
MFC_1 temperature (degC)
MFC_1 volumetic flow (lpm)
MFC_1 mass flow (slpm)
MFC_1 set point (slpm)
V1
V2
V3
V4
Restart FGGA
FGGA Pump | double
CAL MFC_1 Set Value | double

## Ozone
ASRL6::INSTR
Monitoring data | Data Type | Logging Required | Control Data | Data Type | Logging Required | cRIO
o3 conc |  | Yes | Ozone Zeroing (V6) 10mins -> off | Boolean | Yes | NI 9472 - DO4
bench temp |  | Yes | Zero Record Request | Boolean | No
cell a intensity |  | Yes
cell b intensity |  | Yes
pressure |  | Yes
flow a |  | Yes
flow b |  | Yes
flags |  | Yes
Zeroing Record (date, time, zero mean value) |  | No
Log Data Format
Date & Time
CPU Load
MFM
PTP Sync
Flight Number
o3 conc
bench temp
cell a intensity
cell b intensity
pressure
flow a
flow b
flags
Ozone Zeroing (V6)

## 2B Ozone
ASRL2::INSTR
Monitoring data | Data Type | Logging Required | Control Data | Data Type | Logging Required
oz_ppbv |  | Yes | 2B Mode (Ambient or Cabin) | String | Yes
cell_temp |  | Yes | Simulate 3 hour Average | double | No
cell_press |  | Yes | Simulate concentration | I32 | No
vol_flo_r |  | Yes
see UDP for in cabin vs ambient |  | Yes
rolling averages |  | No
Log Data Format
Date
Time
oz_ppbv
cell_temp
cell_press
vol_flo_r
2B Mode

## SO2
ASRL2::INSTR
Monitoring data | Data Type | Logging Required | Control Data | Data Type | Logging Required | cRIO
SO2 conc |  | Yes | SO2 flush (V9) |  |  | NI 9472 - DO7
pmt voltage |  | Yes | SO2 Cal (V9) & MFC_2 set from 0 to user changeable value -> on for X secs -> (V8) on for X secs -> all off & MFC_2 set to 0 | Boolean | Yes | NI 9472 - DO7 -> DO6
lamp voltage |  | Yes | SO2 Zeroing (V6) 10mins -> off | Boolean | Yes | NI 9472 - DO4
lamp int |  | Yes | Zero Record Request | Boolean | No
internal temp |  | Yes | CAL MFC_2 SLPM Value | double |  | RS232 | ASRL4::INSTR
react temp |  | Yes | Sample MFC_3 SLPM Value | double | Yes | RS232 | ASRL5::INSTR
pres |  | Yes
flow |  | No | when wait on wheel is ON, disable V6
flags |  | Yes
rolling averages |  | No
Zeroing Record (date, time, zero mean value) |  | No
MFC_2 absolute pressure (PSIA) | double | Yes
MFC_2 temperature (degC) | double | Yes
MFC_2 volumetic flow (lpm) | double | Yes
MFC_2 mass flow (slpm) | double | Yes
MFC_2 set point (slpm) | double | Yes
MFC_3 absolute pressure (PSIA) | double | Yes
MFC_3 temperature (degC) | double | Yes
MFC_3 volumetic flow (lpm) | double | Yes
MFC_3 mass flow (slpm) | double | Yes
MFC_3 set point (slpm) | double | Yes
V6 | boolean | Yes
V8 | boolean | Yes
V9 | boolean | Yes
MFC 2 Set Value | double | Yes
MFC 3 Set Value | double | Yes
Log Data Format
Date
Time
SO2 conc
pmt voltage
lamp voltage
lamp int
internal temp
react temp
pres
flags
10 mins SO2 averaging
8 hr SO2 averaging
10 mins STEL limit | boolean
8 hr TWA Limit | boolean
MFC_2 absolute pressure (PSIA)
MFC_2 temperature (degC)
MFC_2 volumetic flow (lpm)
MFC_2 mass flow (slpm)
MFC_2 set point (slpm)
MFC_3 absolute pressure (PSIA)
MFC_3 temperature (degC)
MFC_3 volumetic flow (lpm)
MFC_3 mass flow (slpm)
MFC_3 set point (slpm)
V6
V8
V9
MFC 2 Set Value
MFC 3 Set Value

## CalDeck_Control_Requirements
CalDeck valve controls requirements
Caldeck valve control by 24VDC supply; 2000mA max
ev7 - 1.2A surge - needs triggering first to cope with surge
ev7 - 290mA continuous after surge
ev6 - 980mA continuous (no surge)
ev1 - ev4 650mA continuous
ev8 - ev9 325mA continuous
FGGA cal | CO vaccum | SO2 cal
ev1 - ev4 | ev7 | ev8 - ev9
650 mA | 290 mA | 325 mA | 1265 mA
FGGA cal | CO vaccum | o3 / so2 zeroing
ev1 - ev4 | ev7 | ev6
650 mA | 290 mA | 980 mA | 1920 mA | 1900 mA measured
Note: ev6 and ev8 - ev9 MUST BE mutually exclusived in order to operate under 2000mA

## Project Plan
core chemistry development plan
check wiriing | 2015-11-26
Power on test
Rack power on test | 2015-11-30
Hardware Function test
all eValves ON OFF test | 2015-12-01 | Used NI_9472_DO_Test.vi
MFC_1 serial test | 2015-12-01 | Took a few communications restart in order to read parameters
MFC_2 serial test | 2015-12-01 | Took a few communications restart in order to read parameters
MFC_3 serial test
Software instrument comms test
2B ozone | 2015-12-01 | initial serial test without file I/O
CO | 2015-12-01 | initial serial test without file I/O
TEiO3
Tei SO2
FGGA | 2015-12-01 | initial serial test without file I/O
Software user panel design
Software control design
Software control testing
Software status design
Software status testing
Software raw data logging design
Data files storage house keeping
Software raw data logging test